# HPC PLASMA REDUCTION

## PURPOSE

Plasma reduction is indicated for products where volume reduction is required, such as prior to cryopreservation or for recipients of allogeneic HPC products where fluid overload compromises clinical outcome.

## Scope

This procedure describes the plasma reduction process for all haemopoietic products, excluding cord blood and platelet products. Refer to HPC Cord Processing DOC93 and Platelet Processing DOC100.

## Responsibilities

Medical staff are responsible for instructing (Insert manufacturer’s name) of volume reduction requirements for allogeneic HPC products.

## CALCULATIONs

### Plasma Reduction for Cryopreservation

Calculate the volume of plasma to remove:

Target volume of product = (initial WCC  target WCC) x initial volume

Volume plasma to remove = initial volume – target volume.

Target WCC is 600 x109/L (if using 20%DMSO) or 450 x109/L (if using 30% DMSO). Target may be subject to the restrictions of initial WCC and volume.

Volume may be adjusted to fit with cryopreservation bag volume requirements.

## MATERIALS

As listed on the HPC Cryopreservation Form or T Cell Cryopreservation Form.

## Procedure

Plasma reduction is performed on the Sigma centrifuge.

The Thermo Fisher Megafuge 40 may be used as the back-up centrifuge. The procedure is the same as for the Sigma centrifuge.

### Processing using Sigma Centrifuge

- Transfer product into a transfer bag, hermetically sealing line. If there are any visible cell clumps, use an in-line filter. For product volumes >500mL the product may be divided equally into two transfer bags.

- Place the product bag into an outer sterile bag.

- Balance the buckets and load the centrifuge.

- Refer to Equipment Operating Instruction; Sigma Centrifuge.

- Speed setting: 1845rpm (800g)

- Centrifuge for 10min.

- Calculate the volume of plasma to be removed.

- Transfer the double bagged product to the plasma press and hang.

- Insert one transfer pack (300mL or 600mL) with line clamped into the outermost port of the product bag. The attached bag should be labelled as plasma.

- Place the plasma bag on the balance and tare.

- Unclamp plasma bag line and press off plasma to required volume.

- Re-clamp line.

- Carefully remove product bag (with attached plasma bag) from plasma press and place in BSC.

- Insert a SmartSite bag access device into product bag.

- Insert 50mL air using a syringe.

- Mix.

- Repeat process for the second bag (if required).

- Transfer the entire plasma reduced product (from both bags) to the appropriate bag (cryopreservation bag or transfer pack). Measure the transferred volume using a syringe or by weight determination and record.

- Rinse product from each bag and make up to the required volume by allowing plasma to move from the attached plasma bag into the product bag.

- Take required samples (eg for cell counts and, if final product, microbial contamination testing).

## process control

Calculate TNC recovery.

TNC (x1010) = WCC (x109/L) x Volume (L)  10

TNC recovery (%) = TNC (post plasma reduction)  TNC (initial) x 100

TNC recovery should be at least 75%.

If TNC recovery is less than 75%, notify the Production Manager. Cells may be salvaged from the plasma fraction by repeating centrifugation.

## Records

Records of all processing and test results are stored in the (Insert manufacturer’s name) product file.

## forms

HPC Cryopreservation Form FORM76

T Cell Cryopreservation Form FORM106

## Documents

Equipment Operating Instructions DOC5

Sepax 2 Instructions DOC120

HPC Cryopreservation DOC94

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert Date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert Date) | Reviewed (Insert name & description of changes). | (Insert name) | (Insert date) |