# HPC washing

## Policy

For allogeneic HPC products where there is a minor blood group incompatibility between donor and recipient, plasma containing isohaemagglutinins (anti-A or anti-B directed against red cells) is to be removed by washing.

## Purpose

This procedure is for washing the following non-cryopreserved HPC products:

HPC, Marrow (buffy coat enriched)

HPC, Apheresis

## Responsibilities

Medical staff are responsible for ensuring blood group testing is performed for both donor and recipient.

(Insert manufacturer’s name) staff are responsible for

- determining blood group compatibility of allogeneic HPC products.

- processing HPC to remove isohaemagglutinins or density gradient materials.

- any further processing of washed HPC.

## materials

As listed on the HPC Washing Form.

Sepax 2 CS-600.1 kit.

## Procedure

### Preparation of Wash Solution

Prepare 1L of wash solution:

- Transfer 1 x 500ml saline 0.9% and 1 x 500mL Albumex-4 to a 1L Transfer set using the lines of the set.

- Hermetically seal lines.

- Mix.

### Sepax 2

- Follow the procedure Sepax 2 Instructions.

- Two washes should be performed for all non-cryopreserved HPC products that have a minor mismatch.

- If the product volume is >220mL, reduce the volume to less than 220mL prior to commencing the SmartWash Protocol (below). Refer to HPC Plasma Reduction.

- If the product volume is <220mL, proceed directly to the SmartWash Protocol with CS-600.1 kit.

- The collection bag is the input bag.

- Use an appropriate volume transfer bag for the output bag.

- Set the final volume to approximate the initial volume.

- Perform one wash.

- Perform a second wash using the output bag from the first wash as the input bag.

- Use a new transfer bag as the output bag.

- Record the volume.

- Sample the product for TNC count and, if required, a CD34 cell count. An anti-red cell titre is not required, provided 2 washes have been performed.

- Sample for microbial contamination testing, unless proceeding to cryopreservation.

- CFU-GM or viability testing may be performed, as appropriate.

- Insert a SmartSite Add On Access device into the transfer pack for grafts to be immediately infused.

- Complete HPC Washing Form.

## Records

Records of all processing and test results are stored in the (Insert manufacturer’s name) product file.

## documents

HPC Plasma Reduction DOC98

Sepax 2 Instructions DOC120

Sepax: HPC Wash VAL56 rev2

## Forms

HPC Washing Form FORM108

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert Date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert Date) | Reviewed (Insert name & description of changes). | (Insert name) | (Insert date) |