# flow Cytometer maintenance

## purpose

The BD FACSCanto II Flow Cytometer requires little maintenance, however to ensure continued cytometer reliability and optimal functioning, regular basic preventative maintenance must be performed.

## Scheduled Maintenance

Maintenance procedures are summarised in the table below.

Refer to Flow Cytometer Start-up and Shut-down DOC135 for the daily maintenance procedures.

NOTE: Only filtered milli-Q water should be used.

| ### Procedure | ### Description | ### Frequency |
| --- | --- | --- |
| Clean flow cell | Manual clean of the flow cell and SIT with BD FACSFlow solution, then rinse with shutdown solution | ### Daily and as required (sticky dyes or poor optical performance). |
| ### Decontaminating the SIT | Cleans the SIT. Prevents blockage. | ### Daily and as required. |
| ### Purge fluid filters | Removes air from all fluid filters, ensuring they will not dry out | ### Weekly |
| Prepare fresh FACSClean and water tubes | For cleaning flow cell and SIT. | ### Weekly |
| Clean sheath and shutdown fluid cubitainers | Empty and rinse cubitainers with water. Place in wash up room for cleaning. Replace with clean cubitainers. | ### Monthly |
| Decontaminate fluidics (Long Clean) | ### Decontaminates internal sheath path with BD FACSClean solution, then rinse with shutdown solution. Takes approx 75 min. | ### Monthly and before service calls |
| ### Replace air filter | ### Ensures proper cytometer performance | ### Every 6 months by BD. |
| ### Replace fluid filters | ### Keep fluids free of particles | Every 6-12 months, or if increased debris in plots, or leakage under filters. |
| ### Preventative maintenance |  | Every 6 months by BD. |

### Purging the Fluidics Filters (on the fluidics cart)

- Ensure that a Fluidics start up has been performed as part of the daily start up.  Refer to Flow Cytometer Start-up and Shut-down

- Prime the fluidics lines. Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Prime After Tank Refill”.

- Select all the checkboxes for the cubitainers.

- When the completion message appears, click “OK”.

- Put on a pair of gloves and place an absorbent mat beneath the filters to absorb any fluid leakage.

- Purge air from all fluid filters one at a time.

- Gently move the filter around until the air bubble is near the top bleeder valve. Loosen the top bleeder valve of the filter counter-clockwise, but do not completely unscrew the valve as it will come off.

- Wait for fluid to seep out until the bubble disappears.

- Close the bleeder valve by turning it clockwise.

- Repeat steps 7 through 9 with the other two filters.

### Cleaning the Flow Cell

The flow cell must be cleaned daily at shutdown and as required, especially if sticky cells and/or sticky dyes (PI, 7-AAD) are run, or if poor optical performance indicates additional cleaning is required.

- Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Clean Flow Cell”

- When prompted, install a flow tube containing 3mL of BD FACSClean solution onto the SIT and click on “OK”.

- A progress message appears during cleaning of the flow cell which takes approximately 20 seconds.  When the tube has finished running, a completion message appears.

- Repeat steps 1 to 3 twice with the cleaning solution.

- Remove the tube from the SIT.

- Click “OK” to close the completion dialogue box.

- Perform steps 1 to 3 with water three times to remove the cleaning solution from the line and flow cell.

- To clean the SIT, click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “SIT Flush”

- When the SIT cleaning finishes, click “OK” to close the dialog box.

- The cytometer is now ready for use.

### Decontaminating the Fluidics System (Long Clean)

- Check all fluid levels.

- Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Long Clean”

- A confirmation dialog box appears.

- Click “OK” to continue. The process cannot be aborted once the Long Clean has begun.

- Click “OK” to proceed.

If the completion message has not appeared after 90 minutes, verify that there are no error messages in the Status tab of the Cytometer window.

If the cleaning mode fails, refer to the Fluidics Cart Troubleshooting section in this document.

- Select to Shutdown or continue.

  - To shutdown, exit the software and turn off the power to the cytometer

  - To continue, click on “Cytometer, then “Fluidics StartUp” and run the quality control beads.

- If controls fail after a Long Clean, repeat Fluidics Startup until setup passes.

### Replacing the Fluidics Filters

The fluidics cart should be on during this procedure.

- Put on a pair of gloves and place an absorbent mat beneath the filters to absorb any fluid leakage.

- Remove the filter by pressing the metal tabs on each quick-disconnect coupling.

- Position the new filter and connect the couplings. Record the replacement date on the filter.

- Loosen the top bleeder valve on the top of the filter counter-clockwise.

- Wait for fluid to seep out until the bubble disappears.

- Close the bleeder valve by turning it clockwise.

- Repeat the procedure for the two other filters to be replaced.

## Unscheduled Maintenance

Perform these maintenance procedures as indicated in the table:

| ### Procedure | ### Description | ### Frequency |
| --- | --- | --- |
| ### Replace cubitainers |  | ### As required |
| ### Prime fluid lines | ### Removes air from fluid lines | ### Whenever a fluidics line is disconnected to change a cubitainer or for maintenance |
| ### Remove an air lock | Removes air from filter and restores flow of fluid | ### No events in acquiring data, or ### Fluidics are not functioning properly (flow cell will not fill, or there is backflush into sample tubes), or Flow cell and plenum are empty, or Purging the filter is not resolved |
| ### De-gas flow cell | Removes bubbles from flow cell | ### As required during daily start up. |
| ### Bubble filter purge | Removes air from bubble filter | ### If fluidics run dry, or when CV is poor |
| ### Bubble filter purge and De-gas flow cell | Removes bubbles from the bubble filter and the flow cell | ### If air is suspected, or fluidics run dry, or when CV is poor |
| Decontaminate the fluidics system for storage | Cleans out the fluidics lines with BD FACSClean solution and then fills with shutdown solution | ### Before long term storage |
| Replace the Bal seal | ### Replaces worn Bal seal | ### When the seal becomes worn or cracked. The sample tube will not be pressurised. Replaced by BD Biosciences service representative |
| ### Reconnect fluidics cart tubing |  | ### As required |
| ### Replace fluidics level sensor | ### Removes faulty fluid level sensor | ### Error message reports software is unable to get a fluid level, or ### If suspected faulty sensor (first try rinsing it with water), or ### When instructed by BD |
| ### Replace fluidics cart fuses | ### Cart will not function because of power surge or other electrical event. | Replaced by BD |
| ### Reset system circuit breaker switch | ### Restores power to the cytometer and the fluidics cart. | When there is a power surge |

### Priming Fluidics Line

- Ensure that a Fluidics startup has been performed as part of the daily start up.

- Prime the fluidics lines. Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Prime After Tank Refill”.

- Select the checkbox for the cubitainer changed or all the checkboxes if required.  Click “OK”.

- When the completion message appears, click “OK” to close.

### Changing a Cubitainer

The fluidics cubitainers sit on the cart and each cubitainer has its own colour coded port.

- When changing a cubitainer, ensure that the cytometer is not acquiring data.

- Detach the sensor and fluid line from the cart.

- Pull the sensor straight out.

- Press the metal clip of the fluid line.

- Unscrew the cap of the cubitainer.

- Remove the cap and sensor assembly and place aside.

- Place new/clean cubitainer onto the fluidics cart matching the label on the container to the port.

- Replace the cap assembly and tighten until it is fully closed.

- Reattach the sensor and fluid lines to the cart.

- To attach the sensor line, gently rotate it until the connection aligns, and then push in.

- To attach the fluid line, push the coupling into the port until it clicks.

- Prime the fluidics.  See Priming Fluidics Line procedure in this document.

### Removing an Air Lock

- Put on a pair of gloves and place an absorbent mat beneath the air-locked filter(s) to absorb any fluid leakage.

- Remove the filter by pressing the metal tabs.

- Install bypass tubing in place of the filter.

- Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Prime After Tank Refill”.

- Select the checkbox that corresponds to the filter that you have bypassed.  Click “OK”.

- Ensure that fluid is moving through the bypass tubing.

- Remove the bypass tubing when priming finishes.

- Reattach the filter to the fluidics cart.

- Repeat the Priming Fluidics Line procedure of this document.

- Open the top bleeder valve of the filter and wait for fluid to seep out until.  Close the valve.

- If no fluid seeps out, repeat steps 9 and 10.

### Purging the Bubble Filter

- Click on “Cytometer” dropdown menu and select “Cleaning Modes”, then “Bubble Filter Purge and Degas Flow Cell”.

- Wait while the purge finishes.

- When the completion message appears, click “OK” to close.

- Repeat steps 1 to 3 until bubble-free liquid enters the flow cell from the bubble filter.

### Reconnecting the Fluidics Cart Tubing

- Cytometer ports and their corresponding fluidics cart ports are coloured coded.

- Check and reattach fluidics cart tubing if they become detached, in particular the fluid (B) (blue), air out (clear) and waste (orange) ports.

NOTE: The fluidics cart power cord is plugged into the cytometer and NOT into a wall outlet.  This ensures proper electrical grounding and protects against electrical shock or damage to the cytometer.

### Resetting the System Circuit Breaker Switches

The switches must be in the “ON” position at all times.

- The system circuit breaker on the cytometer is a black switch and is located on the left side of the instrument next to the power in plug.

- The circuit breaker on the fluidics cart is a green switch and is located on the left side of the cart next to the blue “fluid (B)” port.

## Troubleshooting

Refer to procedures in this document for details for recommended solutions

| Observation | Possible Causes | Solution |
| --- | --- | --- |
| Flow cell will not fill | - Air in bubble filter - Fluidics cart power - No pressure in the plenum - Sheath line disconnected - Sheath cubitainer empty - Air in the sheath filter (fluidics cart) - Air lock in sheath filter (fluidics cart) - No air pressure in the fluidics cart | - Purge the bubble filter - Turn on the power to the fluidics cart - Check air supply connections & tubing for kinks - Check sheath line connection & tubing for kinks - Replace sheath cubitainer - Purge air from the filter - Remove air lock - Switch off auxiliary air |
| Fluid backfill into sample tube | - Cracked or defective tube - Bal seal worn or improperly installed - Air lock in sheath filter - Tube partially removed from SIT after acquisition | - Transfer sample to another tube (only if non-BD Trucount tube) - Call BD Biosciences - Remove air lock - Continue to remove tube |
| Low laser power indication | - Flow cell access door open - Laser power output is below requirement | - Close door completely - Call BD Biosciences |
| Backflushing of sheath fluid into sample tube | - Aspirator arm bumped or moved to side, but tube was not removed | - Repeat tube removal sequence and perform a manual SIT flush |
| Fluidics pressure errors | - Air lock in filter | - Check the filters in the fluidics cart Verify bottom bleeder valve is tightened If no fluid seeps out, perform Removing an Air Lock |
| Cytometer on, no response to software commands | - Bad keyboard or mouse connection - Communication failure between computer and cytometer | - Check connections to computer - Turn off the computer and cytometer Reseat Ethernet cable, located next to power cord on right side of cytometer Turn cytometer on, then the computer |
| QC fails after a long clean | - Residual BD FACSClean solution in lines | - Run Fluidics Startup to flush the system with sheath fluid Repeat until QC passes |
| Tube does not load, or sample tube does not fit snugly on SIT | - Improper tube used - Defective or cracked tube - Bal seal worn or improperly installed | - Use only BD tubes - Transfer sample to another tube (only if non-BD Trucount tube). Ensure using appropriate tubes - Call BD Biosciences |
| Cytometer and fluidics cart will not turn on | - Power cord disconnected - Cytometer circuit breaker tripped - Blown fuse in cytometer controller | - Reconnect the power cord to the wall socket or cytometer - Switch on the cytometer circuit breaker, followed by the cytometer power if off. - Turn off cytometer main power and call BD |
| Fluid leaking from SIT or aspirator arm | - Interior valve failure or leak - Waste line to fluidics cart disconnected | - Turn off cytometer power and clean up the liquid. Call BD Biosciences - Turn off cytometer power and clean up the liquid. Check that waste lines are securely plugged in Turn on the cytometer power |
| Liquid leakage around cytometer base | - Interior valve failure | - Turn off cytometer power and clean up the liquid. Call BD Biosciences |
| Aspiration arm locked, cannot unload tube | - Software not responding - Pressure loss | - DO NOT use force to remove tube from the SIT Press Ctrl+Alt+Delete to open Task Manager Select End Task to close BD FACSDiva software. Restart software and unload the tube - Verify that the fluidics cart is powered on and that the air pressure hoses are properly connected |
| Fluidics error after replacing sheath fluid | - Status change not yet detected by software | - Wait until software detects status change before starting acquisition (up to 30 secs) |
| Water leakage around fluidics cart base | - Normal condensation overflow from pressure relief valve - Humidity - Bleeder valve on fluidics cart filter open - Broken fluid line | - Turn off cytometer power and clean up the water - Empty the condensation trap in the fluidics cart - Turn off cytometer power and clean up the liquid. Check and close all bleeder valves. Purge the fluidics filters - Call BD Biosciences |
| Fluidics cart will not power on, cytometer on | - Cart circuit breaker off - Auxiliary air switch is on (should be off) - Power cord from fluidics cart to cytometer is disconnected - Fuse blown | - Switch on fluidics cart circuit breaker - Toggle the auxiliary air supply off. When on, the cart’s own air pumps turn off - Connect both ends of the cord - Call BD Biosciences |

## Records

Record maintenance on the Flow Cytometer Checklist FORM156

Store the completed forms in the Flow Cytometer maintenance records file.

## References

- FACSCanto II v2.4 Instructions for Use

- BD FACSCanto II Operator Course Workbook

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert date) | Revised by (Insert name & description of changes) | (Insert name) | (Insert date) |