# temperature mapping

## Purpose

Temperature mapping is performed in spaces where a controlled temperature environment is required, such as incubators and storage areas.

Mapping may be required for the qualification (OQ or PQ) of new equipment.

## equipment

Temperature Data Logger laptop provided by (Insert manufacturer’s name) technical service provider.

## Procedure

### Set-up

- Clean and take laptop, probes, and external power source into the clean room facility.

- Plug the external power source in and turn the power on. External power supply is required as the battery life is insufficient for our requirements (20-24hrs).

- Plug the probes into the USB port located on the back of the laptop.

- Turn the laptop on.  Password is (Insert password if required)

- Select the Temperature Logger icon.  The programme will open.

### To Record

- Select State Control. A drop screen will open.

Select Template. A template files screen will show.

Select the csv template.

Select return.

- Place probes into equipment to be monitored.  The 8 probes are labelled as channels 1-8.

Ensure that the correct channel is in the correct position to the template information.

Channel 1 – Top left rear

Channel 2 – Top centre

Channel 3 – Top right front

Channel 4 – Middle left centre

Channel 5 – Middle front

Channel 6 – Middle right rear

Channel 7 – Bottom left front

Channel 8 – Bottom right rear

- Select State Control.

- Select Record. Do not choose monitoring, as this will not save the data.

- The computer will then prompt a file name.

Enter the equipment name and allocated equipment number (e.g. Inc 32).

Select enter.

- The template is set so that readings are taken every 300secs (5mins).

- The sample time should now be counting down.

- On the main screen you can adjust the temperature units to make it easier to read the graph. For example, if you are measuring 37oC, adjust the Y axis so that the lower limit is 30oC and the upper limit is 45oC.

- At the top off the graph change the window width to 120mins.

- The directory that this will be saved to is automatic, but still check that the Study Directory shows C:\Temperature Logger Data.

- Leave to record for 20-24 hours.

### To Exit

- Select State Control.

- Select Exit.

### Data Extraction

Technical staff may assist with data extraction.

- Insert a USB into the Laptop.

- Raw Data

- Select the Temperature Logger icon.

- Check the Study Directory path for saved data.

- Click on the Start menu.

- Select My Computer.

- Select Directory Path.

- Select saved file and copy to USB (e.g. Inc 32).

- Graph

- Select State control.

- Select Review.

- Select appropriate file.

- A box with the appropriate details on will appear. Close this box.

- On the laptop, use the print screen button.

- Open up the paint option through the Start button.

- Paste (Ctrl V).

- Save as JPG format to the USB.

An example of a graph is printed on the next page. (Insert image of example graph on the next page)

## Records

Save the USB files in (Insert file path here).

## Review

- Results are to be reviewed by the Production Manager/Scientist in Charge and Quality Manager/Quality Coordinator.

- A summary of the results is to be included in the relevant equipment change control (CM) or validation study, with reference to the location of the raw data files.

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert date) | Revised by (Insert name & description of changes) | (Insert name) | (Insert date) |

(Insert image of example graph here)