# MSC viability by flow cytometry

### (BD FACSCanto II Flow Cytometer)

## Principle

7-Amino-Actinomycin D (7AAD) is a fluorescent probe which binds to DNA. It is excluded by viable cells due to their intact cellular membranes. Cells which are dying or completely non-viable have permeable cell membranes which allow the dye to enter the cell and bind to the DNA. As a result, non-viable cells are 7AAD positive.

7AAD fluoresces red (emission spectra 635-675nm) and can be detected by the PCP5.5 detector. It has some spectral overlap with FITC.

## Purpose

To assess the percentage of viable CD90+ Mesenchymal Stromal Cells post cryopreservation, using the BD FACSCanto II Flow Cytometer.

## reagents

FITC Mouse Anti-Human CD90 (BD #555595)

7AAD (BD # 559925)

Flow Buffer (1% BSA in PBS)

Plasma-Lyte

## equipment

BD FACSCanto II Flow Cytometer

Vortex

Pipettes

## Flow Cytometer set-up

- Ensure daily set-up has been performed. Refer to Flow Cytometer Start-up and Shut-down.

- Check quality control results: CS & T Beads and BD FACS 7-Color Setup Beads.

- Open the BD FACS DIVA Software program.

- Log in:

User name: (Insert username)

Password: (refer to Scientist-In-Charge)

- In the browser, highlight the MSC folder and right click on new experiment.

- In the experiment templates dialog box, open the experiment: MSC viability.

- Rename the MSC viability template as: MSC viability_Product#_Lot#.

- In the experiment hierarchy rename the Specimen as: Product#_Lot#.

## Cell Preparation

### Thawing

Remove the sample vial from the cryo storage tank and place into a 37°C water bath until just thawed.

### Staining

- Transfer 5uL of 7AAD and 2.5uL of FITC CD90 into a flow tube.

- Transfer a volume containing approx 2-3 x105cells from the sample into an eppendorf tube. For example, 50uL from a sample with a cell concentration of 5x106 cells/mL.

- Dilute the cell suspension in equal parts with Plasma-Lyte, up to a maximum volume of 100uL.

- Add all the diluted cell solution to the flow tube.

- Wrap the tube in foil (to protect from light) and incubate at room temperature for 10min.

### Measurement

- Add 500uL of flow buffer to the stained cell suspension.

- Immediately run sample on the flow cytometer.

## Result Analysis

### Optimisation

- In the forward scatter and side scatter plot, position the P1 gate to exclude any cell debris and so give clearly defined cell populations in the CD90 + 7AAD plot.

- Position the quadrant gates to optimally capture the live and dead CD90+ cell populations.

- Position the P2 gate in the histogram to separate the total live and dead cell populations.

- Record 20000 events in P1 gate

### Reporting

- Refer to statistics analysis table.

- Record results for the % Parent result in Q4.

- Record result on the MSC Flow Cytometry Form.

- Print a PDF copy of results.

- Save an electronic copy in (Insert file location)

## Records

Hard copy result printouts are stored in the relevant Product file.

Electronic copies are saved in (Insert file location)

## Documents

Flow Cytometer Start-up and Shut-down DOC135

Flow Buffer DOC144

MSC Flow Cytometry Form FORM125

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert Date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert Date) | Reviewed (Insert name & description of changes). | (Insert name) | (Insert date) |