# haemopoietic VIABILITY ASSay

### (BD FACSCanto II Flow Cytometer)

## Principle

7-Amino-Actinomycin D (7AAD) is a fluorescent probe which binds to DNA. It is excluded by viable cells due to their intact cellular membranes. Cells which are dying or completely non-viable have permeable cell membranes which allow the dye to enter the cell and bind to the DNA. As a result non-viable cells are 7AAD positive. The 7-AAD fluorescence is detected in the far red range of the spectrum (655 nm long-pass filter). It has some spectral overlap with FITC and PE, requiring correct compensation settings to be established.

Viability assessment can be performed on any cryopreserved product. The assessment includes:

- percentage of viable CD34+ stem cells or viable CD3+ T cells.

- percentage of viable CD45+ cells.

## purpose and scope

To assess percentage of viable CD34+ cells or CD3+ cells, post cryopreservation, using the BD FACSCanto II Cytometer.

## Responsibilities

(Insert manufacturer’s name) staff members are responsible for performing the flow cytometer daily set up and quality control, routine maintenance and performing the assay.

## REAGENTS

Dilution solution: 0.9% Saline/Albumex 4 (50:50)

CD45-FITC (Immunotech clone J33 - A07782)

CD34-PE (Immunotech clone 581 - A07776)

CD3-PE (Immunotech clone UCHT1 – A07747)

7-AAD Viability Dye (BD Pharmingen 559925)

Flow buffer: PBS with 1% BSA

Pharmlyse red cell lysis buffer (BD 555899)

## EQUIPMENT

BD FACSCanto II Flow Cytometer

Vortex mixer

Pipettes

## sample

Thawed sample from a cryopreserved haemopoietic product e.g. HPC(A), HPC(M), HPC(CB) or T cells.

## PROCEDURE

- If the WCC is greater than 20 x109/L, dilute sample with the saline/Albumex dilution solution. Record the dilution on the Haemopoietic Viability Form.

- Set up one tube with:

- 10 μL of CD45-FITC

- 10 μL CD34-PE (or CD3-PE if assaying T cells)

- 5 μL of 7-AAD

- Add 50 μL of sample (or sample dilution).

- Vortex and incubate for 10 minutes at room temperature.

- At the end of incubation add 500 μL of flow buffer and put through the flow cytometer immediately.

## FLOW ANALYSIS

- Ensure daily set-up has been performed. Refer to Flow Cytometer Start-up and Shut-down and Flow Cytometer Quality Control.

- Check quality control results: CS & T Beads and BD FACS 7-Color Setup Beads.

- Double click on BD FACSDiva Software to open program.

- Open the BD FACSDiva software and log in

User name: (Insert user name)

Password: (refer to Scientist-In-Charge)

- Select “Keep Current Settings” when the CST Mismatch Window appears.

- In the Browser, HPC Viabilities folder and double click to open.

- Highlight the current “month & year” folder and right click to create New Experiment.

- Select the most recent appropriate experiment.

- Select the number of copies required (i.e. the number of viabilities to be measured). Click OK.

- In the Browser, highlight the imported experiment and right click to rename to include the Product Number.

- Click on Specimen Name. Rename as “Product Number and donor’s surname (e.g. 131313 SMITH)

- In the Browser, open the test by clicking on the + (plus) symbol. Click on the arrow. The arrow will turn green to start.

- In the Cytometer Window:

Click on the Laser tab. Check that the FSC Area Scaling is set to 0.65.

Click on the Threshold tab. Check that the FSC parameter value is set to 50,000.

- At the end of incubation, gently vortex the tube and load firmly onto the cytometer.

- Click on acquire data.

Select Flow Rate = medium

NOTE:	Events To Display = 5000

Stopping Gate is set to collect All Events.

Events To Record = 100,000

- Adjust the FS vs SS gate to exclude debris/red cells in bottom left hand corner.

- Adjust the CD45 vs SS gate to include all white cells and use it to exclude any debris.

- Click on Record Data.

- If the number of CD34+ cells is less than 200 events at completion, record more data. Select “Append” from the Warning window.

- The CD34 vs SS gate must include all stem cells and exclude as many dead neutrophils as possible.

- The CD45 vs SS gate on total CD34+ cells is used to exclude any lymphocytes with high CD45.

- Review graphics and re-gate if required. Print results.

- Refer to Flow Cytometer Start-up and Shut-down and Flow Cytometer Quality Control at completion of the assay.

## RESULT processing

- Non-viable cells will show a definite positive 7AAD peak.

- The CD45+ viability region should be placed to include all 7AAD negative events. The CD45+ Viability % refers to the % of all CD45+ cells (i.e. white cells) considered to be viable.

- The CD34+ viability region should be placed to include all 7AAD negative events. The CD34 Viability % refers to the % of all CD34+ events (i.e. stem cells) considered to be viable.

- CD3+ viability assessment is similar to the CD34+ gating strategy except that T cells have higher CD45 expression than stem cells.

## result record

- Results are recorded on the instrument print-out.

- Record the TNC (CD45) and CD34 or CD3 viability results on the Haemopoietic Viability Form.

- Attach the printout to the form and forward both to the Scientist in Charge for review.

- Store the completed form, with print-out attached, in the product file.

## form

Haemopoietic Viability Form FORM135

## Reference

Quantitation of CD34 positive cells – single platform method.

## Procedure History

| Revision | Date | Modification | Approved By | ##### Implemented |
| --- | --- | --- | --- | --- |
| 1.0 | (Insert Date) | Originally drafted by (Insert name) | (Insert name) | (Insert date) |
| 2.0 | (Insert Date) | Reviewed (Insert name & description of changes). | (Insert name) | (Insert date) |